A Microbiologic Comparison of Acellular Dermal Matrices as an Aseptic Reconstructive Material

Shaun D. Mendenhall MD, Katherine E. McKenna BS, Tim Daugherty MS, Nicole M. Cosenza MS, Ryan Schmucker MD, Joel Reichensperger BS, Janak Koirala MD MPH, Paul S. Cederna MD, Michael W. Neumeister MD

Background: Over the last 10 years, the use of acellular dermal matrix (ADM) in breast reconstructive surgery has gained popularity. Unfortunately, 3 recent meta-analyses have demonstrated an increased risk of infection when used for breast reconstruction (1-3). This may be due to the fact that some ADM products are actually not sterile, but instead are “aseptically processed.”

Methods: In order to test the sterility of ADM products, five separate 2x4cm samples of 14 different brands of ADM (Table 1) were sterilely cut into 1x1 cm pieces and placed in liquid culture media for aerobic/anaerobic bacteria, acid-fast bacilli (AFB), and fungi. Standard culture media was incubated for 3 weeks and AFB for 6 weeks. The Biomerieux Vitek 2 system was used to identify organisms from positive cultures. Separate samples of the ADMs were fixed, processed for paraffin embedding, and used for fluorescent in situ hybridization (FISH) using a universal bacterial DNA probe EUB338 to detect any presence of bacterial DNA on the ADMs. FISH slides were evaluated with confocal microscopy.  Differences were tested with the Mann-Whitney U test with p<0.05 considered significant.

Results: The following ADMs had positive cultures (cxs): AlloDerm, AlloDerm RTU, Permacol, DermaMatrix, XCM biologic, Flex HD, and SurgiMend (all were either Bacillus sp., Staph Warneri, or Staph Epidermidis, 1-2 CFU only). Of the positive cxs, 4 were from terminally sterilized ADMs, and 3 were from aseptically processed ADMs. Results of FISH demonstrated traces of bacterial DNA on all matrices except for Integra (Table 1, Figure 1). Number of bacteria per high power field (HPF) on FISH ranged from 0 (Integra) to 13 (AlloDerm) with an average of 2.3. There were more bacteria per HPF in the aseptically processed group compared to the sterile group, although this did not reach significance (3.6 vs 1.9, p= 0.09).

Conclusion: Standard culture techniques of both sterile and aseptically processed ADMs yielded an equal amount of positive cultures, all of which were standard skin flora and were likely contaminants from the culture process. FISH analysis demonstrated evidence of prior bacterial contamination on all ADMs with the exception of Integra. Whether these findings are correlated with clinical infections remains to be studied.

References:

- Kim JY, Davila AA, Persing S, Connor CM, Jovanovic B, Khan SA, Fine N, Rawlani V. A meta-analysis of human acellular dermis and submuscular tissue expander breast reconstruction. Plast Reconstr Surg. 2012 Jan;129(1):28-41.

- Ho G, Nguyen TJ, Shahabi A, Hwang BH, Chan LS, Wong AK. A systematic review and meta-analysis of complications associated with acellular dermal matrix-assisted breast reconstruction. Ann Plast Surg. 2012 Apr;68(4):346-56.

- Hoppe IC, Yueh JH, Wei CH, Ahuja NK, Patel PP, Datiashvili RO. Complications following expander/implant breast reconstruction utilizing acellular dermal matrix: a systematic review and meta-analysis. Eplasty. 2011;11:e40.

- Disclosures/Financial Support:

- Funded by grants from the Memorial Medical Center Foundation and the SIU Dept. of Surgery.

- The authors have no financial interest in any of the products mentioned in the study.  Product support only was provided by Lifenet, Musculoskeletal Transplant Foundation, Synthes, TEI Biosciences, Promethian Life Sciences, and Covidien.